## Slide 1: Commissioning plan for Phase 1
Seiji Kawamura and Osamu Miyakawa
with the inputs from Uchiyama, Michimura, and others
JGW-G1707278-v1

## Slide 2: Objectives and scope of commissioning
According to the roadmap established by SEO, the commissioning team should execute the commissioning activities.
The commissioning activities are to coordinate the various tasks of each subsystem whenever necessary to accomplish the successful operation of the fully-equipped detector with the target sensitivity and duty factor.
Independent installation/integration activities of each subsystem are not included in the scope of commissioning.

## Slide 3: Goal for phase 1
Strict requirement: operation of the cryogenic 3km Michelson interferometer by Mar. 30, 2018
  Mar. 31 reserved for a test run
However, we should complete the phase 1 operation as soon as possible.

## Slide 4: Stages in Phase 1
There are four stages in Phase 1:
Stage 1 (Oct. 2 - End of Oct.)
  Alignment up to BS (inclusive)
Stage 2 (Dec.)
  Alignment of ETMy
Stage 3 (Feb.)
  Locking of MI with Cryogenic ETMy with Calibrated sensitivity curve
Stage 4 (Mar.)
  Locking of MI in vacuum with cryogenic ETMx/y with data stored
Test run (One day before cooling Xend; exact period:TBD)
  Characterization of various systems for Phase 2
(No PRMI operation after test run)

## Slide 5: Stage 1
Initial status
  Beam from IMMT2 roughly aligned to PRM and BS
  PRM, PR2, PR3, and BS installed and integrated

## Slide 6: Rough Schedule
 | 10/2-10/6 | 10/10-10/13 | 10/16-10/20 | 10/23-10/27 | 10/30-11/2 | 11/6-11/10 | 11/13-11/17 | 11/20-11/24 | 11/27-12/1
PRM, PR2, BS integration |  |  |  |  |  |  |  |  | 
Preparation of EXA |  |  |  |  |  |  |  |  | 
Preparation of EYA |  |  |  |  |  |  |  |  | 
Preparation of IXA |  |  |  |  |  |  |  |  | 
Preparation of IYA |  |  |  |  |  |  |  |  | 
Placement of beam tubes, etc. |  |  |  |  |  |  | TBD? |  | 
Preparation of IOO |  |  |  |  |  |  |  |  | 
Alignment of IMMT2, PR2 |  |  |  |  |  |  |  |  | 
Alignment of PR3 |  |  |  |  |  |  |  |  | 
Alignment of BS |  |  |  |  |  |  |  |  | 
Measurement of beam position at IXC, IYC (*1) |  |  |  | Tentative |  |  |  |  | 
Measurement of BS/PR relative position to tanks (*2) |  |  |  |  |  |  | Tentative |  | 

## Slide 7: Laser hazard area
 | 10/2-10/6 | 10/10-10/13 | 10/16-10/20 | 10/23-10/27 | 10/30-11/2 | 11/6-11/10 | 11/13-11/17 | 11/20-11-24 | 11/27-12/1
Upstream of IFI |  |  |  |  | TBD |  |  |  | 
Center area, the 1st floor |  |  |  |  | Tentative |  |  |  | 
Center area, the 2nd floor |  |  |  |  |  |  |  |  | 
Storage and machine room (Kikai-shitsu) | Always laser hazard-free area with safety panels |  |  |  |  |  |  |  | 
EXA |  |  | Hazard-free? |  |  |  |  |  | 
EYA |  |  |  | Hazard-free? |  |  |  |  | 
EXC |  |  |  |  |  |  |  |  | 
EYC |  |  |  |  |  |  |  |  | 
End area, the 2nd floor |  |  |  |  |  |  |  |  | 
Arm tunnel |  |  |  |  |  |  |  |  | 

## Slide 8: Laser hazard area
Center area (Kawamura)
  Place safety panels in front of the Kikai-shitsu and place goggles in the front room (10/10)
  Exclude the OMC area from the hazard area between 11/13 and 11/22 while Fujimi-sangyo is cleaning the area
End area
  Insert the cable etc. into the chamber and cover the space with Aluminum foil (MIF) while the laser is blocked (IOO) so that this area can be always hazard-free, Is this possible?

## Slide 9: Integration of PRM, PR2, BS
PR2の調整
  モーターネットワークの調整をする。ローカルには動くようになったので、リモートで動くかどうかのテスト。
PRM
  Oplevがフレームと干渉してしまった。フレームを取り除くのが大変なので、とりあえずPRMを固定したままで進める。リリースできるかどうかは検討。
BSの調整
  回路の確認、修理。ステッパーモーター、GEO phoneなど。

## Slide 10: Preparation of input optics
Send Laser beam to MIF (IOO)
  Increase laser power to appropriate level
  Bellows will not be put until November.
IMMT2のリリース
  ミラーをリリースしたが、チェンバー中心からマイナスY側に2cmくらいずれていたが、Phase1ではこのまま行く。
  ビームが新しい位置になったので、調整し直し。新井君が手伝う。

## Slide 11: Preparation (EXA/EYA)
Miscellaneous preparation (Uchiyama)
  Clean the EXA/EYA area with Korokoro
  Clean the tool
  G-mat placed for EXA; How about EXY?
  No air gun is prepared immediately
Preparation for the beam measurement (MIF)
  EXA: PD and cable completed
  EYA: PD installed; cable underway
  DAQはKAGRA標準のものになりそう。
  クリーンブース内にパイプが設置してあって、ケーブルはそこを通せばいい。パイプ径は100mm。今回何本も引いてしまえばいいかもしれない(宮川、木村)。
  どんな種類か、何本かをあらかじめ確認しておく。
  安全帯を使うなど、安全には十分に気をつける(木村)。

## Slide 12: Preparation (IXA/IYA)
Cleaning of the area (CRY)
Chambers (CRY)
  Expose IXA/IYA to air
Miscellaneous preparation (Uchiyama)
  Place G-mat on the floor
  Clean the tool
  Preparation of laser hazard area and goggles (Kawamura)

## Slide 13: Preparation (IXC/IYC)
Chambers (CRY)
  Not necessary to remove stuff from IXC
  Partially remove stuff from IYC (10/10)
  Cover the side flange with Aluminum foil (?)

## Slide 14: Preparation of Beam tubes, etc.
PR2-PR3 beam tube (VAC)
  PR2-PR3 beam tube installed
  Open the beam tube and cover the space by aluminum foil (10/10)
  No dry air injection
BS assembly frame (VIS)
  BS assembly frame removed
BS-beam tubes (VAC)
  Y:
    The bellows and the beam tube vonnected to the GV
    Open the connection part and cover the space by Aluminum foil (10/16?)
    No dry air injection
  X:
    The beam tube placed in a proper space
    Open the connection part and cover the space by Aluminum foil (10/16?)
    No dry air injection
IFI-IMM, IMM-PRM, PRM-PR3
  TBD (after alignment and measurement; probably November)

## Slide 15: Alignment
Alignment up to virtual positions of ETMx/ETMy
  10/10 – 10/13: Misalign PRM by a known amount (MIF, VIS), Align IMMT2 (MIF, IOO) and PR2 (MIF, VIS)
  10/16 – 10/20: Align PR3 by monitoring beam in IXA and EXA (MIF, VIS)
  10/23 – 10/27: Align BS by monitoring beam in IYA and EYA (MIF, VIS)
  Notes
    Cables for PD can be put through between the tank and flange.
    ビーム位置の測定:
      情報共有が必要、細かい作業手順を確認したい(麻生)。->内山さんが来週くらいに説明。
      ビーム位置をどう定義、特定するのか？(麻生)
    EXAとEYAでのPDの読み出し:
      情報共有が必要、細かい作業手順を確認したい(麻生)。
      ラフなアラインメントとして、まずはインプット側をゲートバルブの窓で見るほうが効率がいい。
Moving of passageway (Uchiyama)
  Remove the passageway for each chamber during the alignment work
  (During the no-passageway period workers should pass by the wall outside the clean booth.)

## Slide 16: Position measurement
Measurement of beam spot in IXC/IYC with respect to chambers (Uchiyama)
  Accuracy: 1mm?
  Flange in the upper part may or may not be there.
  Moving of passageway
    Remove the passageway for each chamber during the alignment work
    (During the no-passageway period workers should pass by the wall outside the clean booth.)
Measurement of the relative position of BS/PR with respect to chambers (VIS)
  No beam condition is necessary.

## Slide 17: Vacuum check
Probably difficult

## Slide 18: Activities before stage 2
Anything?

## Slide 19: Stage 2
Initial status
  Beam from BS to virtual positions of ETMx/ETMy finely aligned
  ETMy installed and integrated in the air at room temperature

## Slide 20: Activities
Alignment of ETMy
  Align ETMy to IFI (MIF, DGS, VIS, CRY, IOO)
  Lead MI-REFL beam outside the chamber (MIF, IOO)
Installation and alignment of MI-REFL output optics
  Install and align MI-REFL output optics using the beam reflected by ETMy (MIF)
Bellows connection and evacuation of the EYC/EYA (CRY)

## Slide 21: Stage 3
Initial status
  MI aligned except ETMx
  ETMx installed in the air at room temperature
  EYC evacuated

## Slide 22: Activities
RT model for LSC and calibration (DGS, MIF, CAL)
Global control between LSC and VIS models (DGS, MIF, VIS)
Alignment, locking, optimization, characterization of MI
  Align ETMy to IFI (MIF, DGS, VIS, CRY, IOO)
  Make a fine alignment of MI (MIF)
  Lock MI (MIF)
  Optimize and characterize MI (MIF)
Installation and alignment of MI-AS output optics (MIF, VIS)
Cooling-down of ETMy (CRY)
  鏡角度の追跡 シフト？
Calibration (CAL)

## Slide 23: Stage 4
Initial status
  MI locked
  ETMy cooled down

## Slide 24: Activities
Bellows connection and evacuation of the EXC/EXA and cooling-down of ETMx (VAC, CRY)
Bellows connection and evacuation of the whole vacuum system in the center area (VAC)
Re-alignment, re-locking, optimization, characterization of MI (MIF)
Guardian scripts for PSL/IMC/MI (DGS, IOO, MIF)
Data pipe line (DAS)
Physical environment monitors (DetChar)
Selection of monitoring signals (DGS)

## Slide 25: Expected final status
The whole vacuum chamber evacuated
ETMx/ETMy cooled down
MIF re-aligned, locked, optimized
Automatic operation for MI
Detchar system
Various signals integrated into data acquisition system
Various signals monitored at control room

## Slide 26: Laser hazard area
The laser hazard area should be set for the minimum area not to disturb necessary installation/integration work.

## Slide 27: Organization of commissioning team
Organization
  Leader: Kawamura
  Sub-leader: Miyakawa
With the help of
  Scheduler: Uchiyama
  MIF chief: Michimura
  Each subsystem chief/subchief
The commissioning team is a subset of SEO.

## Slide 28: Role of commissioning team
The team establishes the detailed schedule following the overall schedule established by SEO and scheduler.
The team coordinates work of each subsystem to follow the detailed schedule.
At least one of the team members stays at the site (inside the tunnel or at the data acquisition building) everyday during the commissioning period.
When a possible delay of the schedule is anticipated, the team tries to manipulate the order of daily work, etc. to meet the schedule if possible. If not, the team reports to the SEO and asks for the remedy.
The MIF chief helps the team by suggesting the optimized procedure of the commissioning task from the standpoint of the interferometer operation.

## Slide 29: Meeting
The commissioning leader runs a weekly remotely-accessible meeting with the help of sub-leaders between 11:00 and 12:00 on Tuesday after the “Kotei-kaigi” to discuss activities on the previous week and on the coming weeks.
  The team encourages people off site to attend the weekly meeting remotely to provide their inputs.
  Japanese is used for a while.
  Wiki is prepared.
  The document is uploaded to JGWDoc.
  The mailing list is being created. (at some point the list will be reset.)
The leader runs a daily remotely-accessible meeting between 9:00 and 9:10 at the beginning of the “Chorei” to discuss activities on the previous day and on the day .
  The team encourages people off site to attend the daily meeting remotely to provide their inputs.
  eZuce is used.
  Japanese is used for a while.
  Wiki is prepared.
  If we find an issue, we discuss it after the Chorei or so.

## Slide 30: Remote operation (under discussion)
In principle the remote operation from anywhere other than Kamioka (KAGRA site and control room) is prohibited.
If absolutely necessary, remote operation is possible under the following conditions:
  The remote operation should be approved by the commissioning team in advance.
  The corresponding researcher should be present at Kamioka.
  The remote operation should be mentioned in the daily meeting in Chorei.
The remote viewing is OK.

## Slide 31: TBD
We will talk with each subsystem group to optimize the detailed schedule.
  When should the GIF data be implemented in the digital system.
  What kind of evaluation and measurement should be done at each stage?
  When should the bellows in the central area be connected. When should the whole vacuum system in the center area be evacuated? Before or after the first half of the test run?
How long should the test run last?